## How to Use
NFPA 99 RISK ASSESSMENT TOOL
From riskpublishing.com ¦ Structured on NFPA 99 Chapter 4 (4.1 risk categories, 4.2 risk assessment) and the ASHE Systems / Equipment / Emergency Management worksheet layout ¦ Free to adapt
1. Read the Category Key tab. NFPA 99 assigns a category by the worst outcome of a FAILURE of the system or equipment: Category 1 major injury or death, Category 2 minor injury, Category 3 discomfort, Category 4 no impact on patient care.
2. Fill the Systems tab room by room: list each critical room or area and the systems it depends on (medical gas, vacuum, electrical, HVAC, plumbing, IT, nurse call). Enter the worst outcome code (1-4) if that system failed in that room; the sheet returns the category and the highest category per room.
3. Fill the Equipment tab from your medical and building equipment inventory. Enter the failure-outcome code per item; the sheet returns the category and the inspection interval you set on the Category Key tab.
4. Fill the Emergency Management tab with the hazards from your hazard vulnerability analysis and the NFPA 99 Chapter 12 elements each one triggers.
5. Record the assessment team, date, and approver on the Sign-off tab. CMS and The Joint Commission expect the assessment to be documented, defendable, and repeated when the building, systems, or use change.
LEGEND: Blue font = your inputs. Black font = calculated, do not overwrite. Row 2 of each worksheet is a worked example: replace it with your own first entry.
This tool supports the risk assessment NFPA 99 4.2 requires; it does not replace the code text, the authority having jurisdiction, or a qualified engineer. CMS enforces the 2012 edition; The Joint Commission surveys against it under EC.02.05.01.

## Category Key
Code | Category | Failure of the system or equipment is likely to cause | Typical examples | Suggested inspection interval (edit)
1 | Category 1 | Major injury or death to patients, staff, or visitors | Medical gas to ICU and ORs; essential electrical system life-safety and critical branches; ventilators; anesthesia machines | Quarterly
2 | Category 2 | Minor injury to patients, staff, or visitors | Medical gas to general patient rooms; patient monitors on stable patients; nurse call | Semi-annual
3 | Category 3 | Patient discomfort but not injury | HVAC in a waiting area; non-critical lighting; television systems | Annual
4 | Category 4 | No impact on patient care | Administrative office power; landscape irrigation; decorative lighting | Annual or per manufacturer
Source: NFPA 99 Health Care Facilities Code, section 4.1. Category is a property of the failure outcome in that room, not of the equipment type; the same oxygen outlet can be Category 1 in an ICU and Category 2 in an exam room.

## Systems
Room or area | Occupancy / use | Medical gas outcome code (1-4) | Medical vacuum outcome code (1-4) | Electrical outcome code (1-4) | HVAC outcome code (1-4) | Plumbing / water outcome code (1-4) | IT / network outcome code (1-4) | Nurse call outcome code (1-4) | Highest risk | Room category | Notes and existing mitigation
ICU Bed 12 | Inpatient, ventilated patients | 1 | 1 | 1 | 2 | 3 | 2 | 2 |  |  | Bulk oxygen with reserve; EES critical branch; auxiliary source connection added 2026 per NFPA 99-2024 5.1.3.3

## Equipment
Asset ID | Equipment | Location | System it belongs to | Failure outcome code (1-4) | Category | Inspection interval | Last inspected | Next due | Responsible | Notes
EQ-0001 | Oxygen concentrator, portable | Resident room 214 | Medical gas (portable) | 1 |  |  | 2026-06-01 |  | Facilities | Category 1 where used for continuous supply; smoking prohibited in room; concentrator on tested circuit, filters and tubing inspected
EQ-0002
EQ-0003
EQ-0004
EQ-0005
EQ-0006
EQ-0007
EQ-0008
EQ-0009
EQ-0010
EQ-0011
EQ-0012
EQ-0013
EQ-0014
EQ-0015
EQ-0016
EQ-0017
EQ-0018
EQ-0019
EQ-0020
EQ-0021
EQ-0022
EQ-0023
EQ-0024
EQ-0025
EQ-0026
EQ-0027
EQ-0028
EQ-0029
EQ-0030

## Emergency Management
Hazard (from HVA) | Probability (1-3) | Impact (1-3) | Preparedness (1-3, 3 = poor) | Priority score | NFPA 99 Chapter 12 elements triggered | Utility or system dependency | Owner | Exercise or review date
Loss of bulk medical oxygen supply | 2 | 3 | 2 |  | 12.5.3.3 utilities; 12.5.3.3.5 medical gas; auxiliary source connection (2024 ed.) | Bulk oxygen, reserve, auxiliary connection, cylinder cache | Facilities director | 2026-11-15

## Sign-off
Item | Entry
Facility and buildings covered
Assessment date
NFPA 99 edition applied | 2012 (CMS) with 2024 references noted
Team: facilities
Team: clinical / nursing
Team: biomedical engineering
Team: safety officer / emergency management
Approver (name, title)
Next scheduled review
Triggers for interim review | Renovation or change of use; new Category 1 equipment; utility outage or near miss; survey finding; code edition adoption